## signals
signal | unit | description
c_bag_top_weight_act | g | measured bag weight (filtered)
c_bag_top_weight_act_raw | g | measured bag weight
c_bag_rig_weight_act | g | measured bag weight (filtered)
c_bag_rig_weight_act_raw | g | measured bag weight
c_bag_mid_weight_act | g | measured bag weight (filtered)
c_bag_mid_weight_act_raw | g | measured bag weight
c_bag_lef_weight_act | g | measured bag weight (filtered)
c_bag_lef_weight_act_raw | g | measured bag weight
c_bag_top_weight_max | g
c_bag_rig_weight_max | g
c_bag_mid_weight_max | g
c_bag_lef_weight_max | g
c_bag_top_weight_stable | g
c_bag_rig_weight_stable | g
c_bag_mid_weight_stable | g
c_bag_lef_weight_stable | g
c_bag_top_rem_time | s | bag remaining time until empty
c_bag_rig_rem_time | s | bag remaining time until empty
c_bag_mid_rem_time | s | bag remaining time until empty
c_bag_lef_rem_time | s | bag remaining time until empty
c_syr_rem_time | s | syringe remaining time until empty
c_per_test_rem_time | s | remaining time until the periodic test (24h)
c_kit_rem_time | s | kit remaining time until kit change
p_per_test_rem_time | s | remaining time until the periodic test (24h)
p_kit_rem_time | s | kit remaining time until kit change
c_battery_voltage_level | V | battery voltage level
c_on_battery_time | s
c_press_ap_act | mmHg | measured arterial pressure (filtered)
c_press_ap_act_raw | mmHg | measured arterial pressure
p_press_ap_maxw_act | mmHg | arterial pressure window maximum (actual)
c_press_ap_maxw_act | mmHg | arterial pressure window maximum (actual)
p_press_ap_minw_act | mmHg | arterial pressure window minimum (actual)
c_press_ap_minw_act | mmHg | arterial pressure window minimum (actual)
c_press_ap_vp_diff | mmHg
c_press_ep_act | mmHg | measured effluent pressure (filtered)
c_press_ep_act_raw | mmHg | measured effluent pressure
c_press_fp_act | mmHg | measured filter pressure (filtered)
c_press_fp_act_raw | mmHg | measured filter pressure
c_press_sp_act | mmHg | measured solution pressure (filtered)
c_press_sp_act_raw | mmHg | measured solution pressure
c_press_vp_act | mmHg | measured venous pressure (filtered)
c_press_vp_act_raw | mmHg | measured venous pressure
c_pss_3vp | V
c_pss_5vp | V
c_pump_bs_bl_current_act | rpm
c_pump_bs_bl_flow_act | mL/min | blood side blood pump actual flow
c_pump_bs_bl_rot_act | rpm | blood side blood pump actual rotation
c_pump_bs_top_current_act | rpm
c_pump_bs_top_flow_act | mL/h | blood side top pump actual flow
c_pump_bs_top_rot_act | rpm | blood side top pump actual rotation
c_pump_fs_mid_current_act | rpm
c_pump_fs_mid_flow_act | mL/h | fluid side middle pump actual flow
c_pump_fs_mid_rot_act | rpm | fluid side middle pump actual rotation
c_pump_fs_bot_current_act | rpm
c_pump_fs_bot_flow_act | mL/h | fluid side bottom pump actual flow
c_pump_fs_bot_rot_act | rpm | fluid side bottom pump actual rotation
c_pump_syr_flow_act | mL/h | syringe pump actual flow
c_pump_fs_top_current_act | rpm
c_pump_fs_top_flow_act | mL/h | fluid side top pump actual flow
c_pump_fs_top_rot_act | rpm | fluid side top pump actual rotation
c_pump_ca_stop_time | s | syringe pump calcium stop time
c_sadc_air_vol_ul | uL | accumulated air bubble volume by citrate safety air detector (micro litre)
c_sadc_cont_air_vol_ul | uL | continuous air volume by citrate safety air detector (micro litre)
c_sadv_air_vol_ul | uL | accumulated air bubble volume by venous safety air detector (micro litre)
c_sadv_cont_air_vol_ul | uL | continuous air volume by venous safety air detector (micro litre)
c_temp_out_act | C | measured output fluid temperature (filtered)
c_temp_out_act_raw | C | measured output fluid temperature
c_temp_plate_act | C | measured warmer plate temperature (filtered)
c_temp_plate_act_raw | C | measured warmer plate temperature
p_bag_top_weight_act | g | measured bag weight (filtered)
p_bag_top_weight_act_raw | g | measured bag weight
p_bag_rig_weight_act | g | measured bag weight (filtered)
p_bag_rig_weight_act_raw | g | measured bag weight
p_bag_mid_weight_act | g | measured bag weight (filtered)
p_bag_mid_weight_act_raw | g | measured bag weight
p_bag_lef_weight_act | g | measured bag weight (filtered)
p_bag_lef_weight_act_raw | g | measured bag weight
p_bag_top_weight_stable | g
p_bag_rig_weight_stable | g
p_bag_mid_weight_stable | g
p_bag_lef_weight_stable | g
p_css_12v | V
p_css_5v | V
p_css_vsys | V
p_press_ap_act | mmHg | measured arterial pressure (filtered)
p_press_ap_act_raw | mmHg | measured arterial pressure
p_press_ap_vp_diff | mmHg
p_press_ep_act | mmHg | measured effluent pressure (filtered)
p_press_ep_act_raw | mmHg | measured effluent pressure
p_press_fp_act | mmHg | measured filter pressure (filtered)
p_press_fp_act_raw | mmHg | measured filter pressure
p_press_fp_maxw_act | mmHg | upper limit of filter pressure
c_press_fp_maxw_act | mmHg | upper limit of filter pressure
p_press_fp_minw_act | mmHg | lower limit of filter pressure
c_press_fp_minw_act | mmHg | lower limit of filter pressure
p_press_sp_act | mmHg | measured solution pressure (filtered)
p_press_sp_act_raw | mmHg | measured solution pressure
p_press_vp_act | mmHg | measured venous pressure (filtered)
p_press_vp_act_raw | mmHg | measured solution pressure
p_press_vp_maxw_act | mmHg | upper limit of venous pressure
c_press_vp_maxw_act | mmHg | upper limit of venous pressure
p_press_vp_minw_act | mmHg | lower limit of venous pressure
c_press_vp_minw_act | mmHg | lower limit of venous pressure
p_pump_bs_bl_flow_act | mL/min | blood side blood pump actual flow
p_pump_bs_bl_rot_act | rpm | blood side blood pump actual rotation
p_pump_ci_flow_req | mL/h
p_pump_bs_top_flow_act | mL/h | blood side top pump actual flow
p_pump_bs_top_rot_act | rpm | blood side top pump actual rotation
p_pump_fs_mid_flow_act | mL/h | fluid side middle pump actual flow
p_pump_fs_mid_rot_act | rpm | fluid side middle pump actual rotation
p_pump_fs_bot_flow_act | mL/h | fluid side bottom pump actual flow
p_pump_fs_bot_rot_act | rpm | fluid side bottom pump actual rotation
p_pump_ca_flow_req | mL/h
p_pump_syr_flow_act | mL/h | syringe pump actual flow
p_pump_fs_top_flow_act | mL/h | fluid side top pump actual flow
p_pump_fs_top_rot_act | rpm | fluid side top pump actual rotation
p_sadc_air_vol_ul | uL | accumulated air bubble volume by citrate safety air detector (micro litre)
p_sadc_cont_air_vol_ul | uL | continuous air volume by citrate safety air detector (micro litre)
p_sadv_air_vol_ul | uL | accumulated air bubble volume by venous safety air detector (micro litre)
p_sadv_cont_air_vol_ul | uL | continuous air volume by venous safety air detector (micro litre)
p_supercap_3vp | V
p_temp_out_act | C | measured output fluid temperature (filtered)
p_temp_out_act_raw | C | measured output fluid temperature
p_temp_plate_act | C | measured warmer plate temperature (filtered)
p_temp_plate_act_raw | C | measured warmer plate temperature
p_temp_inner | C
p_temp_battery | C
c_acc_bl_vol_act | L | accumulated (treated) blood volume
c_acc_ca_vol_act | mL | accumulated calcium volume (during treatment)
c_acc_ci_vol_act | L | accumulated citrate volume
c_acc_di_vol_act | L | accumulated dialysate volume
c_acc_ef_vol_act | L | accumulated effluent volume
c_acc_fl_vol_act | mL
c_acc_hep_vol_act | mL | accumulated heparin volume
c_acc_net_rem_vol_act | mL | accumulated net fluid removal volume
c_acc_pl_su_vol_act | mL | accumulated plasma substitution volume
c_acc_pl_vol_act | mL | accumulated plasma volume
c_acc_su_pos_vol_act | L | accumulated post substitution volume
c_acc_su_pre_vol_act | L | accumulated pre substitution volume
c_acc_therapy_time_act | s | elapsed therapy time
c_bag_waste_vol | mL
c_bag_test_vol | mL
p_bag_test_vol | mL
c_bag_test_weight | g
p_bag_test_weight | g
c_test_vol_3 | g
p_test_vol_3 | g
c_test_vol_4 | g
p_test_vol_4 | g
c_filtration_ratio | % | fluid/blood filtration ratio calculated by the machine
c_net_rem_flow_act | mL/h | actual net fluid removal flow
c_net_rem_vol_act | mL | actual net fluid removal volume
c_net_rem_vol_est | mL | estimated net fluid removal volume
c_press_fdp_act | mmHg | actual filter pressure drop
c_press_tmp_act | mmHg | actual transmembrane pressure
c_blood_vol_act | mL | actual delivered blood volume
c_blood_flow_act | mL/min | actual blood flow
c_pump_bl_vol_act | L
c_pump_ca_vol_act | mL
c_pump_ci_vol_act | mL
c_pump_ci_vol_req | mL
c_pump_di_vol_act | mL
c_pump_ef_vol_act | mL
c_pump_fl_vol_act | mL
c_pump_hep_bol_vol_act | mL | actual delivered heparin bolus volume
c_pump_syr_vol_comp | mL
c_syringe_rem_vol_act | mL | syringe remaining volume
c_pump_pl_su_vol_act | mL
c_pump_pl_vol_act | mL
c_pump_su_p_vol_act | mL
c_pump_su_g_vol_act | mL
c_pump_di_vol_req | mL
c_pump_su_p_vol_req | mL
c_pump_su_g_vol_req | mL
c_ca_trtd_bl_vol | mL
c_ext_bl_vol_stop | mL
c_ext_bl_vol_ther | mL
d_renal_dose_act | mL/h/kg | renal dose measured by machine
d_target_renal_dose | mL/h/kg | average target renal dose based on set values
d_renal_dose_set | mL/h/kg | current target renal dose based on set values
c_bag_ci_weight_diff | g
c_bag_di_weight_diff | g
c_bag_ef_weight_diff | g
c_bag_su_p_weight_diff | g
c_bag_su_g_weight_diff | g
p_bag_ci_weight_diff | g
p_bag_di_weight_diff | g
p_bag_ef_weight_diff | g
p_bag_su_g_weight_diff | g
p_bag_su_p_weight_diff | g
p_pump_bl_vol_act | L
p_pump_ci_vol_act | mL
p_pump_di_vol_act | mL
p_pump_ef_vol_act | mL
p_pump_su_g_vol_act | mL
p_pump_su_p_vol_act | mL
p_net_rem_vol_act | mL | actual net fluid removal volume
p_net_rem_vol_est | mL
p_press_fdp_act | mmHg | actual filter pressure drop
p_press_tmp_act | mmHg | actual transmembrane pressure
p_pump_ca_vol_act | mL
p_pump_hep_bol_vol_act | mL
p_pump_syr_rot_act | rpm
p_pump_hep_vol_act | mL
d_pump_su_pg_flow_act | mL/h | summed substitution flow in post-post dilution mode
c_kit_bs_volume_ml | mL
p_kit_bs_volume_ml | mL
c_cpat_balance_act | mL
c_bag_rig_weight_net_act | g
c_net_rem_flow_est | mL/h
c_pump_fs_mid_flow_est | mL/h
c_pump_fs_top_flow_est | mL/h
p_therapy_time_act | s
p_therapy_time_res | s
p_pump_fs_mid_flow_est | mL/h
p_net_rem_flow_est | mL/h
g_net_rem_flow_set | mL/h | net fluid removal flow set by the user
g_therapy_time_set | s | therapy time set by the user
g_calcium_ratio_set | mmol/L | calcium ratio set by the user
c_10v | V
c_24v | V
c_vcc | V
p_battery_voltage | V | measured battery voltage
p_battery_current | mA | measured battery current
c_citrate_ratio_act | mmol/L | actual citrate ratio value
c_calcium_ratio_act | mmol/L | actual calcium ratio value
c_calcium_flow_sum_req | mmol/h | required total calcium flow
c_calcium_flow_by_su_p_act | mmol/h | actual calcium flow added by substitution
c_calcium_flow_by_su_p_percent | % | percentage of calcium flow added by substitution
g_citrate_ratio_set | mmol/L | citrate ratio set by the user
c_bag_rig_weight_empty | mL
c_net_rem_flow_set | mL/h
c_pump_fs_mid_flow_set | mL/h
g_bag_rig_weight_net_set | mL
g_syringe_rem_vol_set | mL
g_filtr_ratio_alarm_lim_set | %
g_filtr_ratio_info_lim_set | %
g_pump_bl_flow_set | mL/min | blood flow set by the user
g_pump_di_flow_set | mL/h | dialysate flow set by the user
g_pump_hep_flow_set | mL/h | heparin flow set by the user
g_pump_su_pre_flow_set | mL/h | pre-substitution flow set by the user
g_pump_pl_su_vol_set | mL | plasma sbstitution volume set by the user
g_pl_su_ratio_set | % | plasma filtration ratio set by the user
g_temp_out_set | C | fluid temperature set by the user
g_patient_data_weight_set | kg | patient weight set by the user
g_net_rem_vol_set | mL | net fluid removal set by the user
g_press_ap_max_set | mmHg | maximum arterial pressure limit set by the user
g_press_ap_min_set | mmHg | minimum arterial pressure limit set by the user
g_press_ap_maxw_set | mmHg | max arterial pressure window limit set by the user
g_press_ap_minw_set | mmHg | min arterial pressure window limit set by the user
g_press_ep_min_set | mmHg | minimum effluent pressure limit set by the user
g_press_fdp_max_set | mmHg | max filter pressure drop limit set by the user
g_press_fp_max_set | mmHg | max filter pressure limit set by the user
g_press_fp_maxw_set | mmHg | max filter pressure window limit set by the user
g_press_fp_minw_set | mmHg | min filter pressure window limit set by the user
g_press_tmp_max_set | mmHg | max transmembrane pressure limit set by the user
g_press_vp_maxw_set | mmHg | max venous pressure window limit set by the user
g_press_vp_minw_set | mmHg | min venous pressure window limit set by the user
g_pump_su_pos_flow_set | mL/h | post-substitution flow set by the user
g_pump_hep_bol_vol_set | mL | heparin bolus volume set by the user
g_patient_data_hct_set | % | patient hematocrit value set by the user
g_syringe_filled_vol_set | mL | syringe filled volume set by the user
g_bag_mid_num_set |  | middle bag numbers set by user
c_bag_mid_num_act |  | middle bag numbers estimated by machine
p_pump_bs_bl_tube_const |  | blood side blood pump tube constant
c_filter_factor |  | filter (dialyser) factor measured by machine
c_trmt_main_state |  | control sw main state
c_trmt_sub_state |  | control sw sub state
p_trmt_main_state |  | protective sw main state
p_trmt_sub_state |  | protective sw sub state
g_trmt_main_state_set |  | graphical user interface sw main state
g_trmt_sub_state_set |  | graphical user interface sw sub state
g_ci_bag_conc_set |  | set citrate concentration
g_ca_syr_conc_set |  | set calcium concentration
g_anticoag_mode_set |  | id of the selected anticoagulation mode
g_substitution_mode_set |  | id of the selected substitution mode
g_therapy_mode_set |  | id of the selected therapy type
d_serial_number_to_odi |  | Serial number
d_kit_type_str |  | kit type
g_patient_id_str |  | patient name
g_syringe_type_set_str |  | syringe type
g_bag_mid_num_set |  | middle bag numbers set by user
c_bag_mid_num_act |  | middle bag numbers estimated by machine
p_pump_bs_bl_tube_const |  | blood side blood pump tube constant
c_filter_factor |  | filter (dialyser) factor measured by machine
c_trmt_main_state |  | control sw main state
c_trmt_sub_state |  | control sw sub state
p_trmt_main_state |  | protective sw main state
p_trmt_sub_state |  | protective sw sub state
g_trmt_main_state_set |  | graphical user interface sw main state
g_trmt_sub_state_set |  | graphical user interface sw sub state
g_ci_bag_conc_set |  | set citrate concentration
g_ca_syr_conc_set |  | set calcium concentration
g_anticoag_mode_set |  | id of the selected anticoagulation mode
g_substitution_mode_set |  | id of the selected substitution mode
g_therapy_mode_set |  | id of the selected therapy type
d_serial_number_to_odi |  | Serial number
d_kit_type_str |  | kit type
g_patient_id_str |  | patient name
g_syringe_type_set_str |  | syringe type

## equivalences
signal | value | name
c_trmt_main_state | 0 | Preparation
| 1 | Connect patient
| 2 | Therapy
| 3 | End of therapy
| 4 | Temporarily disconnect
| 5 | Change kit
| 6 | Change therapy
| 7 | Change dilution
| 8 | Change anticoag
c_trmt_sub_state | 0 | Preparation start screen
| 1 | Preparation history view
| 2 | Preparation previous kit removal
| 3 | Preparation previous kit removal stop
| 4 | Preparation previous kit removal unloading
| 5 | Preparation scan kit
| 6 | Preparation select therapy
| 7 | Preparation device test
| 8 | Preparation device test finished
| 9 | Preparation install kit
| 10 | Preparation install kit stop
| 11 | Preparation install kit loading
| 12 | Preparation install kit unloading
| 13 | Preparation install kit ready
| 14 | Preparation install bags
| 15 | Preparation install bags ready
| 16 | Preparation priming
| 17 | Preparation priming stop
| 18 | Preparation priming blood side filling start
| 19 | Preparation priming blood side art line filling
| 20 | Preparation priming blood side ven line filling
| 21 | Preparation priming blood side fpc level adjustment
| 22 | Preparation priming blood side vpc level adjustment
| 23 | Preparation priming blood side filled
| 24 | Preparation priming fluid side filling start
| 25 | Preparation priming fluid side dial line filling
| 26 | Preparation priming fluid side ef line filling
| 27 | Preparation priming fluid side su g pre line filling
| 28 | Preparation priming fluid side su g post line filling
| 29 | Preparation priming fluid side su p main line filling
| 30 | Preparation priming fluid side plasma line filling
| 31 | Preparation priming fluid side flushing line filling
| 32 | Preparation priming fluid side spc level adjustment
| 33 | Preparation priming fluid side bypass line filling
| 34 | Preparation priming fluid side seq ef line filling
| 35 | Preparation priming fluid side epc level adjustment
| 36 | Preparation priming fluid side filled
| 37 | Preparation priming dialyser degassing
| 38 | Preparation priming required rinsing
| 39 | Preparation priming final test
| 40 | Preparation priming final rinsing
| 41 | Preparation priming rinsing finished
| 42 | Preparation priming ready
| 43 | Preparation ready for therapy
| 44 | Preparation optional rinsing
| 45 | Preparation optional rinsing ready
| 46 | Preparation recirculation
| 47 | Preparation confirm
| 48 | Preparation cancel
| 49 | Preparation cancel stop
| 50 | Preparation cancel unloading
| 51 | Preparation leaving
| 52 | Connect patient
| 53 | Connect patient stop
| 54 | Connect patient running ac stop
| 55 | Connect patient running ac running
| 56 | Connect patient ready
| 57 | Therapy main
| 58 | Therapy blood stop
| 59 | Therapy patient care blood circ ac stop
| 60 | Therapy patient care blood circ ac running
| 61 | Therapy blood circ ac stop
| 62 | Therapy blood circ ac running
| 63 | Therapy running ac stop
| 64 | Therapy running ac running
| 65 | Therapy periodic test
| 66 | End of therapy return blood
| 67 | End of therapy return blood stop
| 68 | End of therapy return blood running
| 69 | End of therapy return blood ready
| 70 | End of therapy disconnect all
| 71 | End of therapy disconnect all stop
| 72 | End of therapy disconnect all unloading
| 73 | End of therapy leaving
| 74 | Temporarily disconnect patient return blood
| 75 | Temporarily disconnect patient return blood stop
| 76 | Temporarily disconnect patient return blood running
| 77 | Temporarily disconnect patient return blood ready
| 78 | Temporarily disconnect patient disconnect pat
| 79 | Temporarily disconnect patient recirculation
| 80 | Temporarily disconnect patient recirculation stop
| 81 | Temporarily disconnect patient recirculation running
| 82 | Temporarily disconnect patient confirm
| 83 | Temporarily disconnect patient connect patient
| 84 | Temporarily disconnect patient connect patient stop
| 85 | Temporarily disconnect patient connect patient running ac stop
| 86 | Temporarily disconnect patient connect patient running ac running
| 87 | Temporarily disconnect patient connect patient ready
| 88 | Change therapy select therapy
| 89 | Change therapy connect lines
| 90 | Change therapy confirm
| 91 | Change therapy blood stop
| 92 | Change therapy blood circ ac stop
| 93 | Change therapy blood circ ac running
| 94 | Change dilution select dilution
| 95 | Change dilution connect lines
| 96 | Change dilution confirm
| 97 | Change dilution blood stop
| 98 | Change dilution blood circ ac stop
| 99 | Change dilution blood circ ac running
| 100 | Change kit return blood
| 101 | Change kit return blood stop
| 102 | Change kit return blood running
| 103 | Change kit return blood ready
| 104 | Change kit disconnect kit
| 105 | Change kit disconnect kit stop
| 106 | Change kit disconnect kit unloading
| 107 | Change anticoag select anticoag
| 108 | Change anticoag unload syringe
| 109 | Change anticoag change bags
| 110 | Change anticoag confirm
| 111 | Change anticoag blood stop
| 112 | Change anticoag blood circ ac stop
g_trmt_main_state_set | 0 | Preparation
| 1 | Connect patient
| 2 | Therapy
| 3 | End of therapy
| 4 | Temporarily disconnect
| 5 | Change kit
| 6 | Change therapy
| 7 | Change dilution
| 8 | Change anticoag
g_trmt_sub_state_set | 0 | Preparation start screen
| 1 | Preparation history view
| 2 | Preparation previous kit removal
| 3 | Preparation previous kit removal stop
| 4 | Preparation previous kit removal unloading
| 5 | Preparation scan kit
| 6 | Preparation select therapy
| 7 | Preparation device test
| 8 | Preparation device test finished
| 9 | Preparation install kit
| 10 | Preparation install kit stop
| 11 | Preparation install kit loading
| 12 | Preparation install kit unloading
| 13 | Preparation install kit ready
| 14 | Preparation install bags
| 15 | Preparation install bags ready
| 16 | Preparation priming
| 17 | Preparation priming stop
| 18 | Preparation priming blood side filling start
| 19 | Preparation priming blood side art line filling
| 20 | Preparation priming blood side ven line filling
| 21 | Preparation priming blood side fpc level adjustment
| 22 | Preparation priming blood side vpc level adjustment
| 23 | Preparation priming blood side filled
| 24 | Preparation priming fluid side filling start
| 25 | Preparation priming fluid side dial line filling
| 26 | Preparation priming fluid side ef line filling
| 27 | Preparation priming fluid side su g pre line filling
| 28 | Preparation priming fluid side su g post line filling
| 29 | Preparation priming fluid side su p main line filling
| 30 | Preparation priming fluid side plasma line filling
| 31 | Preparation priming fluid side flushing line filling
| 32 | Preparation priming fluid side spc level adjustment
| 33 | Preparation priming fluid side bypass line filling
| 34 | Preparation priming fluid side seq ef line filling
| 35 | Preparation priming fluid side epc level adjustment
| 36 | Preparation priming fluid side filled
| 37 | Preparation priming dialyser degassing
| 38 | Preparation priming required rinsing
| 39 | Preparation priming final test
| 40 | Preparation priming final rinsing
| 41 | Preparation priming rinsing finished
| 42 | Preparation priming ready
| 43 | Preparation ready for therapy
| 44 | Preparation optional rinsing
| 45 | Preparation optional rinsing ready
| 46 | Preparation recirculation
| 47 | Preparation confirm
| 48 | Preparation cancel
| 49 | Preparation cancel stop
| 50 | Preparation cancel unloading
| 51 | Preparation leaving
| 52 | Connect patient
| 53 | Connect patient stop
| 54 | Connect patient running ac stop
| 55 | Connect patient running ac running
| 56 | Connect patient ready
| 57 | Therapy main
| 58 | Therapy blood stop
| 59 | Therapy patient care blood circ ac stop
| 60 | Therapy patient care blood circ ac running
| 61 | Therapy blood circ ac stop
| 62 | Therapy blood circ ac running
| 63 | Therapy running ac stop
| 64 | Therapy running ac running
| 65 | Therapy periodic test
| 66 | End of therapy return blood
| 67 | End of therapy return blood stop
| 68 | End of therapy return blood running
| 69 | End of therapy return blood ready
| 70 | End of therapy disconnect all
| 71 | End of therapy disconnect all stop
| 72 | End of therapy disconnect all unloading
| 73 | End of therapy leaving
| 74 | Temporarily disconnect patient return blood
| 75 | Temporarily disconnect patient return blood stop
| 76 | Temporarily disconnect patient return blood running
| 77 | Temporarily disconnect patient return blood ready
| 78 | Temporarily disconnect patient disconnect pat
| 79 | Temporarily disconnect patient recirculation
| 80 | Temporarily disconnect patient recirculation stop
| 81 | Temporarily disconnect patient recirculation running
| 82 | Temporarily disconnect patient confirm
| 83 | Temporarily disconnect patient connect patient
| 84 | Temporarily disconnect patient connect patient stop
| 85 | Temporarily disconnect patient connect patient running ac stop
| 86 | Temporarily disconnect patient connect patient running ac running
| 87 | Temporarily disconnect patient connect patient ready
| 88 | Change therapy select therapy
| 89 | Change therapy connect lines
| 90 | Change therapy confirm
| 91 | Change therapy blood stop
| 92 | Change therapy blood circ ac stop
| 93 | Change therapy blood circ ac running
| 94 | Change dilution select dilution
| 95 | Change dilution connect lines
| 96 | Change dilution confirm
| 97 | Change dilution blood stop
| 98 | Change dilution blood circ ac stop
| 99 | Change dilution blood circ ac running
| 100 | Change kit return blood
| 101 | Change kit return blood stop
| 102 | Change kit return blood running
| 103 | Change kit return blood ready
| 104 | Change kit disconnect kit
| 105 | Change kit disconnect kit stop
| 106 | Change kit disconnect kit unloading
| 107 | Change anticoag select anticoag
| 108 | Change anticoag unload syringe
| 109 | Change anticoag change bags
| 110 | Change anticoag confirm
| 111 | Change anticoag blood stop
| 112 | Change anticoag blood circ ac stop
g_anticoag_mode_set | 0 | Not selected
| 1 | None
| 2 | Citrate calcium
| 3 | Heparint
g_substitution_mode_set | 0 | Not selected
| 1 | Post
| 2 | Pre post
| 3 | Pre
| 4 | Post post
| 5 | None
g_therapy_mode_set | 0 | No valid therapy selected
| 1 | Slow Continuous Ultrafiltration
| 2 | Continuous Veno-Venous Hemofiltration
| 3 | Continuous Veno-Venous Hemodialysis
| 4 | Continuous Veno-Venous Hemodiafiltration
| 5 | Hemoperfusion
| 6 | Therapeutic Plasma Exchange
| 7 | Cascade Filtration
| 8 | Plasma adsorption/perfusion
| 9 | Low Blood Volume Slow Continuous Ultrafiltration
| 10 | Low Blood Volume Continuous Veno-Venous Hemofiltration
| 11 | Low Blood Volume Continuous Veno-Venous Hemodialysis
| 12 | Low Blood Volume Continuous Veno-Venous Hemodiafiltration